HARDWARE ARCHITECTURE

Bilateral Pressure-Sensing Mat System

for Post-Operative Rehabilitation Monitoring on Rowing Ergometer

Prepared for: Kashif (PCB Design & Fabrication)

Prepared by: Dean Arvan

Date: March 2026

Revision: 1.0 (Mat-Based Design)

CONFIDENTIAL — Patent Pending

Bilateral Force-Sensing System and Method for Post-Operative Musculoskeletal Rehabilitation Monitoring with Dynamic Baseline Compensation

# 1. System Overview

This document specifies the hardware architecture for a bilateral pressure-sensing mat system designed to mount directly onto a Concept2 rowing ergometer foot stretcher. The system replaces the earlier discrete-FSR insole design with a high-density pressure matrix that provides true 2D pressure surface mapping, dramatically improving center-of-pressure (CoP) accuracy and enabling surface pressure visualization.

## 1.1 Design Rationale

The previous design used 3 discrete FSR sensors per foot (heel, ball, toe) inside a shoe insole. While functional for basic bilateral symmetry detection, this approach has critical limitations for clinical rehabilitation monitoring:

- Spatial resolution: 3 discrete points cannot resolve mediolateral CoP shifts within the foot, missing key ACL compensation patterns.

- Shoe variability: Different footwear creates inconsistent sensor coupling, degrading session-to-session repeatability.

- Calibration fragility: Single-point calibration (standing body weight) becomes inaccurate at forces 2-3x standing load during rowing drive.

The mat-based design solves all three by providing a dense pressure grid on a fixed, controlled surface. The foot stretcher is rigid and the foot position is constrained by the straps, making it an ideal sensing platform.

## 1.2 System Block Diagram

The system consists of two identical pressure mat assemblies (left foot, right foot), connected via flat flex cables to a single electronics enclosure mounted on the ergometer frame.

| LEFT MAT Flex PCB + Velostat 12×28 matrix (336 pts) FFC-40 connector | ELECTRONICS ESP32-S3 MCU 4× CD74HC4067 Mux LiPo + TP4056 BLE Antenna | RIGHT MAT Flex PCB + Velostat 12×28 matrix (336 pts) FFC-40 connector |
| --- | --- | --- |

# 2. Sensor Mat Assembly

Each foot pad is a three-layer sandwich that mounts on a Concept2 foot stretcher. The foot stretcher plate is approximately 120mm wide × 280mm tall per foot, angled at 38-42° from horizontal.

## 2.1 Layer Stack-Up

| Layer | Material | Function | Thickness | Source |
| --- | --- | --- | --- | --- |
| Top Cover | Medical-grade silicone (Dragon Skin 10 or Ecoflex 00-30) | Moisture barrier, non-slip foot contact surface, force distribution | 1.0 – 1.5 mm | Smooth-On |
| Top Flex PCB | Polyimide (Kapton), 0.1mm, ENIG finish, 12 column traces | Column electrodes (vertical strips, 10mm pitch) | 0.1 mm | JLCPCB / PCBWay |
| Sensing Layer | Velostat (3M 1361) or Linqstat piezoresistive film | Variable resistance under pressure (30-50 kΩ unloaded → <1 kΩ at 200N+) | 0.1 mm | 3M / Adafruit / Amazon |
| Bottom Flex PCB | Polyimide, 0.1mm, ENIG finish, 28 row traces | Row electrodes (horizontal strips, 10mm pitch) | 0.1 mm | JLCPCB / PCBWay |
| Bottom Cover | Silicone + adhesive backing or Velcro | Stretcher attachment, moisture barrier | 1.0 – 1.5 mm | Smooth-On + 3M VHB tape |

Total mat thickness: approximately 3.5 – 4.5 mm. This is thin enough to not significantly alter the ergometer foot feel or strap tension.

## 2.2 Flex PCB Specifications

| Parameter | Specification |
| --- | --- |
| Board outline | 120mm × 280mm active area + 20mm tab for FFC connector |
| Substrate | Polyimide (Kapton), 0.1mm (4 mil) thickness |
| Copper weight | 1 oz (35µm) — standard for flex |
| Surface finish | ENIG (Electroless Nickel Immersion Gold) — corrosion-resistant for sweat exposure |
| Layers | Single-layer per PCB (2 separate PCBs: one for rows, one for columns). No vias needed. |
| Trace width | 8mm wide traces, 2mm gap between traces (10mm pitch) |
| Bottom PCB traces | 28 horizontal traces (rows), terminating at 28-pad connector footprint on tab |
| Top PCB traces | 12 vertical traces (columns), terminating at 12-pad connector footprint on tab |
| Connector | 40-pin FFC/FPC connector, 0.5mm pitch, bottom contact (28 rows + 12 columns = 40 signals) |
| Design note | The two flex PCBs are separate parts. They are aligned and bonded with the Velostat sheet sandwiched between them during assembly. Row traces run perpendicular to column traces. |

DESIGN NOTE: Trace Alignment The row and column traces must be precisely aligned so each intersection forms a consistent sensing point. Consider adding alignment holes or fiducial marks at the corners of both flex PCBs. The Velostat sheet is cut to the same active area (120×280mm) and placed between the two PCBs with no adhesive — the silicone encapsulation holds the sandwich together.

## 2.3 Sensing Matrix Geometry

| Parameter | Value |
| --- | --- |
| Rows (horizontal) | 28 (driven sequentially by row mux) |
| Columns (vertical) | 12 (read simultaneously or sequentially by column mux) |
| Sensing points per foot | 336 (28 × 12) |
| Total sensing points (bilateral) | 672 (336 × 2 feet) |
| Spatial resolution | 10mm × 10mm per sensing point (1 cm²) |
| Electrical connections per mat | 40 (28 rows + 12 columns) via single FFC-40 cable |

# 3. Scanning Circuit Design

## 3.1 Operating Principle

The system reads the pressure matrix by sequentially driving each row HIGH and measuring the resulting voltage on each column. At each row-column intersection, the Velostat acts as a variable resistor whose resistance decreases with applied pressure.

## 3.2 Voltage Divider Configuration

Each column line connects to ground through a fixed reference resistor (10 kΩ). When a row is driven to 3.3V, the voltage at each column intersection forms a voltage divider:

V_column = 3.3V × R_ref / (R_ref + R_velostat)

| Condition | R_velostat | V_column | ADC Reading (12-bit) |
| --- | --- | --- | --- |
| No pressure | ~30 – 50 kΩ | ~0.6 – 0.8 V | ~750 – 1000 |
| Light pressure (~20N) | ~10 – 15 kΩ | ~1.3 – 1.7 V | ~1600 – 2100 |
| Medium pressure (~100N) | ~2 – 5 kΩ | ~2.2 – 2.8 V | ~2700 – 3500 |
| Heavy pressure (~300N+) | ~0.5 – 1 kΩ | ~2.9 – 3.1 V | ~3600 – 3900 |

## 3.3 Row Driving Circuit

Rows are driven one at a time using analog multiplexers. A 330Ω series resistor on the mux common pin limits current in case of a short circuit or concentrated point load.

- Row Mux (Left Mat): 2× CD74HC4067 (16-channel each). Mux A handles rows 0-15, Mux B handles rows 16-27. Selected by 4 shared address pins (S0-S3) + individual enable pins (EN_A, EN_B).

- Row Mux (Right Mat): 2× CD74HC4067 (identical configuration). Selected by the same 4 address pins + separate enable pins (EN_C, EN_D).

- Total row mux ICs: 4× CD74HC4067 (2 per mat).

## 3.4 Column Reading Circuit

Columns are read sequentially through a column multiplexer into the ESP32 built-in ADC.

- Column Mux (Left Mat): 1× CD74HC4067 (channels 0-11 used). 12 column signals routed to ESP32 ADC1 channel via common pin.

- Column Mux (Right Mat): 1× CD74HC4067 (channels 0-11 used). 12 column signals routed to ESP32 ADC2 channel via common pin.

- Pull-down resistors: 24× 10 kΩ (12 per mat), connecting each column line to ground.

- Total column mux ICs: 2× CD74HC4067 (1 per mat).

## 3.5 Scanning Timing Budget

| Operation | Time | Notes |
| --- | --- | --- |
| Set row mux address (4 GPIO writes) | ~1 µs | Negligible |
| Velostat settling time (RC constant) | 50 – 100 µs | Wait after row switch |
| Read 12 columns via ADC (sequential) | ~120 µs | ESP32 ADC at ~10µs/read |
| Total per row | ~220 µs |  |
| Total per mat (28 rows) | ~6.2 ms |  |
| Total both mats (sequential scan) | ~12.4 ms | Well within 20ms budget |
| Remaining budget (at 50Hz = 20ms frame) | ~7.6 ms | For CoP computation, BLE TX, biofeedback |

PERFORMANCE NOTE At 50Hz frame rate, the system comfortably scans 672 sensing points (both mats) in 12.4ms, leaving 7.6ms for signal processing and BLE transmission. The system could run at 80Hz (12.5ms frame) and still have margin. Running at 100Hz (10ms frame) would require optimizing the Velostat settling time or accepting slightly noisier readings.

# 4. Main Electronics Board

## 4.1 Bill of Materials

| # | Component | Part Number | Qty | Pkg | Est. Cost | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | ESP32-S3-WROOM-1 Module | ESP32-S3-WROOM-1-N8R2 | 1 | Module | $3.50 | BLE 5.0, dual ADC |
| 2 | 16-Channel Analog Mux | CD74HC4067M | 6 | SOIC-24 | $3.00 | 4 row + 2 column |
| 3 | 10 kΩ Resistor (pull-down) | RC0402FR-0710KL | 24 | 0402 | $0.25 | 12 per mat column |
| 4 | 330Ω Resistor (row current limit) | RC0402FR-07330RL | 4 | 0402 | $0.10 | 1 per row mux output |
| 5 | 40-pin FFC Connector (0.5mm) | FH12-40S-0.5SH | 2 | SMD | $1.00 | Hirose, bottom contact |
| 6 | 40-pin FFC Cable (0.5mm, 150mm) | Custom length | 2 | Cable | $2.00 | Length depends on mounting |
| 7 | LiPo Charge Controller | TP4056 (with DW01A protection) | 1 | SOP-8 | $0.30 | USB-C charge input |
| 8 | 3.3V LDO Regulator | AMS1117-3.3 | 1 | SOT-223 | $0.15 | 500mA output |
| 9 | LiPo Battery | 603048 (or similar) | 1 | Pouch | $4.00 | 3.7V 800mAh |
| 10 | USB-C Connector | USB4105-GF-A | 1 | SMD | $0.50 | Charging + firmware update |
| 11 | Decoupling Capacitors (100nF) | CL05B104KO5NNNC | 10 | 0402 | $0.10 | 1 per IC VCC pin |
| 12 | Slide Switch (SPDT) | SS-12D00 | 1 | THT | $0.20 | Main power on/off |
| 13 | Status LED (RGB common cathode) | LTST-C19FD1WT | 1 | SMD | $0.30 | Power/BLE/charging status |

Estimated electronics BOM cost per unit: $15 – $20 (excluding flex PCBs and Velostat)

Estimated flex PCB cost: $15 – $30 for 5 panels from JLCPCB (per foot, so $30 – $60 for bilateral pair)

Estimated total unit cost: $50 – $85 for a complete bilateral system

## 4.2 ESP32 Pin Assignment

| ESP32 Pin | Signal | Direction | Description |
| --- | --- | --- | --- |
| GPIO 4 | MUX_S0 | Output | Mux address bit 0 (shared, all 6 muxes) |
| GPIO 5 | MUX_S1 | Output | Mux address bit 1 |
| GPIO 6 | MUX_S2 | Output | Mux address bit 2 |
| GPIO 7 | MUX_S3 | Output | Mux address bit 3 |
| GPIO 15 | ROW_EN_A | Output | Left mat row mux A enable (rows 0-15) |
| GPIO 16 | ROW_EN_B | Output | Left mat row mux B enable (rows 16-27) |
| GPIO 17 | ROW_EN_C | Output | Right mat row mux C enable (rows 0-15) |
| GPIO 18 | ROW_EN_D | Output | Right mat row mux D enable (rows 16-27) |
| GPIO 1 (ADC1) | COL_READ_L | ADC Input | Left mat column mux output (analog read) |
| GPIO 2 (ADC1) | COL_READ_R | ADC Input | Right mat column mux output (analog read) |
| GPIO 8 | COL_MUX_EN_L | Output | Left column mux enable |
| GPIO 9 | COL_MUX_EN_R | Output | Right column mux enable |
| GPIO 48 | STATUS_LED | Output | RGB LED data (WS2812 or direct drive) |

Note: GPIO pins 1 and 2 are on ADC1 (not ADC1_CH0). ADC2 cannot be used while WiFi/BLE is active on ESP32-S3, so all analog reads must be on ADC1 channels. The column mux outputs for left and right mats connect to separate ADC1 channels, allowing sequential or interleaved reading.

# 5. Assembly Instructions

- Cut the Velostat sheet to 120mm × 280mm (matching the flex PCB active area). Cut two pieces, one per foot.

- Align the bottom flex PCB (row traces, horizontal) face-up on a clean, flat surface.

- Place the Velostat sheet on top of the bottom PCB, ensuring full coverage of all traces.

- Place the top flex PCB (column traces, vertical) face-down on top of the Velostat, with traces facing the Velostat.

- Verify alignment: row and column traces should be perpendicular. Each intersection should be centered.

- Secure the sandwich temporarily with small clips at the edges.

- Place the sandwich in a silicone mold (or use a dam frame) and pour Dragon Skin 10 silicone to encapsulate. Leave the FFC tab exposed.

- Allow silicone to cure (4-7 hours at room temperature, or 1 hour at 60°C).

- Attach the FFC cable to the connector tab. Apply strain relief (hot glue or silicone bead) at the cable exit point.

- Mount the completed mat on the Concept2 foot stretcher using 3M VHB double-sided tape or Velcro strips.

# 6. Software Interface

The firmware running on the ESP32-S3 performs the matrix scanning loop and outputs per-foot summary data over BLE at approximately 1 stroke per second. The existing rehabilitation software modules receive the same data format as the previous discrete-sensor design, ensuring backward compatibility.

## 6.1 Scanning Loop (Pseudocode)

for each mat in [LEFT, RIGHT]: for row = 0 to 27: enable_row_mux(mat, row) // Set S0-S3, enable correct mux IC delay_microseconds(80) // Wait for Velostat RC settling for col = 0 to 11: set_col_mux_address(col) matrix[mat][row][col] = adc_read(mat_adc_pin) disable_row_mux(mat, row) // Compute per-foot summary from 336-point matrix total_force[mat] = sum(matrix[mat]) cop_x[mat] = weighted_centroid_x(matrix[mat]) // ML position cop_y[mat] = weighted_centroid_y(matrix[mat]) // AP position pressure_map[mat] = downsample(matrix[mat]) // Optional: for surface plots

## 6.2 Output Data Format (Per Frame, BLE)

Each 50Hz frame produces the following summary data for each foot. One BLE notification per stroke (~1Hz) transmits the stroke-averaged values.

| Field | Type | Unit | Description |
| --- | --- | --- | --- |
| total_force | uint16 | ADC units (or calibrated N) | Sum of all 336 sensing points |
| cop_x | int16 | 0.01mm resolution | Mediolateral CoP (signed) |
| cop_y | int16 | 0.01mm resolution | Anteroposterior CoP (signed) |
| heel_zone | uint16 | ADC units | Sum of rows 0-9 (posterior 1/3) |
| ball_zone | uint16 | ADC units | Sum of rows 10-21 (middle 1/3) |
| toe_zone | uint16 | ADC units | Sum of rows 22-27 (anterior 1/3) |
| peak_pressure | uint16 | ADC units | Maximum single-point reading |
| contact_area | uint8 | count (0-336) | Number of points above noise threshold |

The heel/ball/toe zone sums provide backward compatibility with the existing software modules, which expect 3-zone force data per foot. The cop_x and cop_y fields provide the true 2D centroid that the new posturalStability module uses for sway analysis.

# 7. Design Considerations and Open Questions

## 7.1 Velostat Characterization

Velostat has known nonlinearity, hysteresis, and creep (drift under sustained load). For the prototype phase, this is acceptable because the primary metrics are bilateral ratios (LSI, symmetry index) where both mats drift similarly. For production, consider upgrading to Sensitronics pressure-sensitive sheet material, which has better linearity and published force-resistance curves.

## 7.2 Calibration Strategy

The existing bilateral standing calibration routine (patient stands still, known body weight) translates directly to the mat system. Each of the 336 sensing points gets a scale factor derived from its share of the total standing force. Multi-point calibration (applying known weights at specific locations) is recommended for production units.

## 7.3 In-Session Drift Compensation

During the recovery phase of each rowing stroke (when the seat slides forward and foot force drops to near-zero), the firmware can capture a zero-force reference reading for each sensing point. This per-stroke auto-zero compensates for Velostat creep over the session duration. This is a firmware feature, not a hardware change.

## 7.4 Ergometer Compatibility

The Concept2 Model D/E foot stretcher has adjustable heel height (positions 1-6) and nylon straps. The mat must be thin enough to not interfere with strap tension or heel cup. The 3.5-4.5mm total thickness of the silicone-encapsulated mat is comparable to a standard insole and should not affect ergometer feel. The FFC cable should exit from the lateral (outside) edge of each mat to avoid interference with the foot straps.

## 7.5 Future Enhancements

- Balance board mat: Same electronics, smaller mat (e.g., 200mm × 200mm) for standalone postural stability assessments.

- Higher density: Reduce trace pitch from 10mm to 5mm for 4× resolution (1344 points/foot). Requires faster scanning or interleaved reading.

- Sensitronics upgrade: Replace Velostat with calibrated pressure-sensitive film for production units.

- Wireless mat: Embed the ESP32 and battery directly in the mat encapsulation for a fully wireless, zero-cable design.

# 8. Requested Deliverables from Kashif

- Flex PCB layout for the bottom layer (28 row traces) with FFC-40 connector footprint. Gerber files for JLCPCB fabrication.

- Flex PCB layout for the top layer (12 column traces) with FFC-40 connector footprint. Gerber files for JLCPCB fabrication.

- Main electronics board schematic and PCB layout (rigid, 4-layer, approximately 40mm × 60mm) with all BOM components placed. Gerber files for fabrication.

- Schematic review: verify mux addressing, ADC channel assignment, power regulation, and BLE antenna clearance.

- DRC (Design Rule Check) clean pass on all three boards.

- Assembly notes for the flex-Velostat-flex sandwich alignment.

End of Document